## Open Action list
Item # | Date Raised | Action Item | Description of issue | Comments/updates | Responsible | Status
1 | 2022-12-16 | LDL Performance issue | Concerns expressed by a number of GP’s that LDL results > Cholesterol. Because of INAB findings, LDL is not currently being reported & there is a calculation in place as an alternative. Graham has expressed that he feels clinically exposed because of this issue & requires a resolution as soon as possible. Currently reviewing correlation data between old & new method. From provisional data, it looks like the new method may not resolve the bias completely. Still approximately 5% bias but additional data needs to be generated to confirm. It would be helpful to review the validation data to understand how the Atellica LDL performance compared to the Beckman method. The updated Lipid assays (LDL, HDL & Trigs) all need to go live concurrently. This may involve configuring a rule on ADM where a calculated LDL may be required where the rules are violated. | Trevor & Conor to configure trigger for particular samples in ADM. Data for LDL assay, as well as other lipid assays, has been sent to Paul & Martina. Graham to review. Plan to resume running 'old' LDL method with rule in place until a decision is made based on the data. Looks like new method is approx 5-6% different from old method but to be confirmed. 13/01/23 Graham reviewing data, waiting sign off, possibly next week. Trevor and Conor have calculation working on ADM side. 20/01 Rule now in place for fasting & random lipids. If rule is breached (>15%), samples will be sent to the Mater for cholesterol testing. 27/01/23 Conor seeing 1 or 2 per day. Rule is working. Continue to monitor for the moment before closing out. 03/02/23 Plan to move to new assays next week. ADM rules to be moved to new method - TREVOR 10/02/23 Some results are being held in ADM & some in the LIS which is causing a delay in Clinicans seeing results. Trevor to investigate with Martina | Olive/Enda/Graham/Paul/Martina/Trevor
2 | 2022-12-16 | Batch testing - samples missed | Conor noted that x2 batches of LDL samples (x55 samples total) were front loaded onto the Sample Handler and on both occasions, only 18 results were produced – to be investigated by Cruinn | Olive advised that she successfully ran batches of lipid samples across all 3 platforms - To be monitored & tested. 13/01/23 still being monitored, OMG and Conor to do this next week. 20/01 Continues to be monitored - Conor to run dummy batch of samples when possible. 27/01/23 Olive to do this when next on site. 10/02/23 Olive to test when next on site 24/03/23 Batch testing completed & Olive confirmed that all samples were tested. Olive to run through batch testing procedure with RHM staff. | Olive/Conor
3 | 2022-12-16 | Lithium Assay | 10 day TAT from Eurofins is not acceptable. Validation of Advia method is still ongoing. Cruinn will press Siemens for further information on the root cause of the issue. | All data generated has been sent to Paul/Martina for review. 13/01/23 Graham signed off on this and going live on Monday. 20/02 New test name (Xlith) to be added to SMS/Aptio server - MF to complete today. Closed | Emma/Olive/Martina |  | Ongoing
4 | 2022-12-16 | RSM Rack issue | Racks not being recognised & racks getting stuck – Emmett is on site 16/12 to troubleshoot this issue. | Monitor last week, no errors. Issue closed. Closed | Emmet/Michael |  | Completed
5 | 2022-12-16 | Dilution issue | CRP & UCFP automated dilutions are not working correctly. Discussion around whether these should be configured on Atellica Software rather than ADM to make appropriate use of the on-board dilution ring. Need to understand if this is an issue for all dilutions or specific to these 2 assays. | Work completed on diluton hardware on instruments - likely to have resolved issue - needs to be monitored 13/01/23 Neill worked on this during the week and its working as expected now. 20/01 May need to configure for Trig dilutions. HCG initial dilution causing delayfor EPU samples. - need to invesitgate if front loading EPU samples may resolve this 27/01/23 Neill working on analyser so unavailable, to get update next week 03/02/23 Paul to confirm if CRP & UCFP dilutions are working correctly. | Neill
|  |  | Dilution should be completed if the dilution request is on the instrumet. On board dilution is sufficient for up to 15 photometric tests - unlikely to have >15 on UCFP samples. To be checked on CRP samples | Dilutions may need to be configured on Atellica rather than ADM to ensure on board dilution is used rather than needing sample to be retrieved from RSM. 13/01/23 Reference email from NOL sent Tue 10/01/2023 14:22 for details. Closed | Trevor/Olive
6 | 2022-12-16 | EQA Testing | Need to configure the system to run EQA samples on a particular CH930 system on the long line. Neill initially hoped this could be configured through ADM but not successful. Is it possible to set up a panel for a particular instrument – Olive/Enda to investigate. | Olive spoke with UK FAS & this has been configured on ADM in Scottish site - Olive to follow up & get specific configuration details. 13/01/23 Neiling continue working on this, setup configuration and needs to test it. Instruction sent to Conor on how to run this. 20/01 - Configuration seems to be working well except when a reagent is missing from the particular system. Staff will need to know what assays are not available on board. 27/01/23 Lab to define configuration and test. 03/02/23 Lab testing to be completed 10/02/23 Update needed from Conor 24/02/23 Working effectively except if a reagent is missing on board, samples will be routed to different analyser | Lab
7 | 2022-12-16 | Sodium Issue | There was a complaint recently from a GP regarding high sodiums and a positive bias was noted. Enda completed an adjustment on the IMT’s which improved the patient bias but this does not seem evident in the IQC. | All data generated has been sent to Graham for review. 13/01/23 Graham reviewed this data and emailed response today, waiting on further eqa results. Could be the end of the month before we get these results. OMG to review precision at higher end. Paul to send data table to OMG. Compare with Temple St data and Siemens UK. 20/01 Graham to provide feedback on data sent by OMG this week. EQA results are on not aligned with Atellica group - might be helpful to assess EQA data from Temple St & Crumlin - OMH/EM to investigate. 27/01/23 NOKLUS (similar to EQA scheme). Paul needs patient medians from ADM (Trevor) so this data can be exported and sent to NOKLUS for analysis. To be done weekly. MF has sent Trevor details of work to be actioned. 03/02/23 Update needed from Trevor 10/02/23 Trevor to investigate data pull for NOKLUS 24/02/23 Martina to ask Graham for format of data required for NOKLUS submission 03/03/23 Graham to send exampe of file format required for NOKLUS submission 24/03/23 Sample file format provided by Graham. Trevor to review & faciliate file transfer as required. 31/03/23 This item has been included in the IT list provided. | Trevor/Paul /Martina
8 | 2022-12-16 | Direct ISE - Sodium issue | Poor EQA performance on sodium. Data to be provided & shared with Glenbio to investigate. | Emma to follow up on status with Glenbio. 13/01/23 Possibly in wrong user group on NEQAS. Paul following this up. Emma to have a call with Edmond/Glenbio on this also. 20/01 Emma to contact Glenbio again. 27/01/23 Paul has sent data to Glenbio, awaiting feedback. Emma to follow up with Edmond also. 03/02/23 Data with Sarah in Glenbio for reveiw - Emma to follow up 10/02/23 Emma to follow up with Glenbio 24/02/23 Emma to follow up with Glenbio 03/03/23 Have requested an on site meeting with Glenbio to determine an appropriate plan 10/03/23 Rachel from Glenbio expected on site today to troubleshoot issue. 24/03/23 Slow progress investigating the issue by Glenbio. Sodoium performance compares well against the Atellica method. 31/03/23 Requested on site support from Regina in Glenbio to run any requred testing as part of the investigation. | Emma/Glenbio
9 | 06/12/202 | IT Troubleshooting out of hours | Oliver requested that an SOP be put together to troubleshoot IT issues out of hours such as a flowchart on what to test/check | 13/01/23 Oliver/Trevor working on this, should Siemens always be notified if an IT issue arise, PH yes Siemens should be notified. SOP in progress. 20/01 SPO being finalised 27/01/23 ADM went down Wednesday 25/01 night, Oliver got it working. Went down and did not restart. Invetigate can DMS buffer results if ADM is down and then resend results when ADM comes back online. Trevor following this up with Siemens. 03/02/23 Investigation ongoing with Siemens on issue from 25/01; Trevor to check if restart protocol is working correctly. 10/02/23 Oliver to go through SOP with lab staff to ensure they understand the troubleshooting process 24/02/23 SOP to be amended to be simpler. Key staff members to be identified to have access level to reboot servers etc, Peter to inquire with Janice on server responsibilities. 03/03/23 Oliver has amended the SOP, key staff to be trained; Recent IT fail investigated & Oliver identified that at 11.02 an ADM backup kicked in & at 11.30 the syste crashed - timing seeems incorrect so Oliver & Neill/Michael to review. Peter confirmed with Janice that the server is supported by Siemens. 10/03/23 Additional IT downtime on 03/03 - Trevor to provide update. Looks like similar issue to previous downtime. If issue reoccurs, lab to log issue with Siemens, need to identify what internal update is causing the issue. Peter to follow up with Siemens following call. 24/03/23 SOP to be finalised by Trevor & Oliver & key staff then to be trained. Awaiting feedback from National ICT as to whether a snapshot is been taken which may be the cause of the server crashing. 31/03/23 Call with National ICT - they advised that there is no snapshots being performed & advised that the issue is not at VM level. Call scheduled for 04/04 @ 8am to discuss & some configuration changes may be required which will result in a server reboot | Trevor /Oliver/Peter
10 | 2023-12-06 | QC Volume on Sample Handler | Issue ongoing since installation - QC samples are being kicked off SH as the software is detecting no volume, despite there being sufficient volume present (on both line) | 13/01/23 Declan to give update on this next week. To be discussed with Janice/Simens on next Thursdays call. 20/01 Call with Janice is next Thurs 26/01. 27/01/23 To get update next week. 03/02/23 Siemens investigation ongoing 10/03/23 Neill & Declan completed some work on this issue - awaiting the outcome 03/03/23 Neill advised RHM not to overfill QC tubes as this may be causing some of the issue. Neill to advise exact volume & staff to be trained by Martina/Conor/Paul. 10/03/23 Olive calulated exact QC volumes - correct volumes to be aliquoted next week & outcome to be assessed. Never more than 4ml to be aliquoted. 24/03/23 No issues to report from staff with new volumes implemented but needs to be monitored. 31/03/23 Continue to monitor | Olive/Enda
11 | 2023-12-06 | Cals & QC processing | Calibrations & QC testing is not being prioritised during the working day - advised that this will be improved with the upcoming SW update which is on controlled roll-out. Waiting on UK sites to upgrade first | 13/01/23 workarounds not useful. Problem is with a 6 point cal using a lot of carriers. Real fix is pending a s/w upgrade, NOL following this up. PH to follow up with Siemens Healthcheck team. 46% capacity analysers being queried. 20/01 HCS Healthcheck requested - awaiting dates (March at earliest). EC/PH to query this with Siemens next week during US visit. 27/01/23 To get update next week. 03/02/23 PH/EC advised that this issue was discussed while in the US last week. The resolution of this is top priority & is actively being worked on 03/03/23 Peter advised that a fix is now in place for this issue. Beta testing is now beginning & we expect a SW upgrade mid year will resolve this issue. 31/03/23 Martina advised that the lab is adjusting the process to prep all cals the evening before & run them first thing in the mornings which should eliminate/reduce the need to cal during the day | Emma/Peter
12 | 2023-12-06 | Flags for low reagent levels | Advised that users are having trouble identifiying when reagents are low or run out. Tablets are no in use. Too many flags on the instruments to identify reagent supply issues. | Olive/Enda to provide training on tablets. Olive/Enda to look at instrument flags & setting up threshold levels for reagent supplies. DMS also has reagent flags - Neill to show users 13/01/23 OMG looked at tablets and reviewed reagent levels, work still needed on this, Declan to help. User training also required. MF following up on DMS to see if this can be used for reviewing reagent levels. 20/01 Martina advised that tablets do not work well for monitoring & flagging reagent levels - there are too many flags to decifer. EC/PH to provide feedbeck to Siemens next week. MF to investigate DMS reagent flags. 27/01/23 No further update, Olive and Enda to follow up with Siemens. 03/02/23 Expectation that SW 1.27 will improve this. Once installed we will review how reagent levels are flagged & determine if this is effective. 10/03/23 Ticket to be raised | Olive/Enda/Neill/Michael
13 | 2023-12-06 | Atellica Process Manager | To be installed | 13/01/23 Peter followed this up with Siemens/ Justin Jone, waiting feedback. 20/01 Trevor to complete APM training - awaiting training info. Siemens will install APM & suppport will then be provided by Cruinn. 27/01/23 In progress, no further update to report. 24/02/23 Trevor training dates pushed out until April. Once completed, the installation can be scheduled. | Peter/Trevor
14 | 2023-12-06 | QC Review on Atellica SW | x2 levels of QC still cannot be viewed on long line (short line is fine). Olive advised that this is all configured in ADM & to work with Paul to show him how to review the data & generate reports | 13/01/23 Olive and Neill looked at this, problem is on line 1 as more data to review. Fix is to use ADM. Olive following up with Trevor next Wednesday on site. 20/01 Olive & Enda to complete ADM QC training & then train RHM staff on QC review. 27/01/23 To get update next week 03/02/23 Once Olive/Enda have completed their ADM QC training, all staff will be trained 03/03/23 Peter to send screenshots to Siemens Global on this issue 10/03/23 Awaiting feedback from Global; Ticket to be raised 24/03/23 No improvement in QC review. Also noted that there is now an issue excluding data points - to be added to the IT list 31/03/23 Olive raised a ticket regarding the omission of data points - awaiting feedback from Chris Walker | Olive/Enda
15 |  | Technopath IAMQC | Technopath IAMQC to be progressed | 13/01/23 Oliver waiting on feedback from Martina. 20/01 Oliver & Martina to follow up with Technopath as the non-conformance needs to be closed. 27/01/23 No movement on this, non - comformance issue which need to be closed in less than 4 weeks. Patrick in technopath cannot be contacted currently. Paul has aksed can Peter follow up and try and get movement on this issue. 03/02/23 Waiting on feedback from Oliver 10/02/23 Upload from ADM has been configured by Trevor. Trevor & Oliver to work on this & communication file location & access to Technopath 24/02/23 Awaiting confirmation from Tony that file format is correct. 03/03/23 Need confirmation from Technopath that the data is correct; Martina advised that there is some 'cleanup' required on the file pulled from Atellica 24/03/23 Emma to follow up with Technopath as this is not progressing 31/03/23 Discussion with Technopath at Biomedica - mails were in spam folder. This can now progress & Cruinn will monitor to ensure completion. | Technopath/Oliver/Trevor
16 |  | Serum Indices verification | Additional testing required on HIL methods. Biorad material to be ordered by Enda | 13/01/23 Call to be setup with Mullingar and Siemens Global. Emma to setup this call next week. 20/01 Emma to send out dates for conf call with Global to discuss 27/01/23 Set for next Tuesday 31/01/2023 03/02/23 Based on call this week, plan to order in Biorad QC material & Jason to dial in to SRS & review data 10/02/23 Concern that the Biorad QC material may not be beneficial. WIll proceed with testing this & go from there 24/02/23 Order for material placed - long lead time from Fannins. 03/03/23 Biorad material has arrived in Cruinn & will be shipped next week for evaluation. Graham expressed frustration at the HIL index that is provided rather that the definitive result . To be fed back to Siemens. Greg Simpson mentioned there is a site in Holland that is reporting numbers - Emma to investigate at Biomedica. 10/03/23 Biorad material to be run up next week to assess performance. Conor indicated that both index & absolute values are avaialble but ADM configuration would need to be changed. Need to investigate if it is possible to send across both values? Ticket to be raised 24/03/23 Biorad material has not yet been run up; confirmed it was delivered. Follow up meeting scheduled with Global on 30/03. 31/03/23 Call yesterday - After reviewing instrument sata through SRS, Jason confirmed the instruments are performing as expected. Plan to run study to compare sodium performance on lipaemic samples on Atellica CH, Direct ISE & Beckman AU systems (in Crumlin). Also noted that the cutoff for Atellica Lipaemai is 124 & for Beckman AU is 40. This info is to be shared with Jason. | Enda/Martina
17 |  | Ward Sample priorotisation | TAT for Ward samples is same day. However, samples are taking too long. Can samples from a particular location be prioritised for testing? | 13/01/23 Possible use of ASAP option on Aptio SMS. MF to get more infomration on how it operates. Paul and Michael to discuss further. 20/01 MF to configure lane on Aptio GUI & test functionality. 27/01/23 MF/Paul to trial next week 03/02/23 MF/Paul to trial 10/02/23 MF advised that the configuration at the IOM is not useful. Discussed the possibility of changing prioritisation of samples in NetAquire to ASAP. MF to forward email to Martina & Oliver to investigate if this is possible. 24/02/23 Prioritisation of samples on NetAquire list | Trevor/Conor/Michael
18 |  | Sample disposal | It seems that samples that have scheduled/rerun tests are being disposed of from the RSM. Can a rule be configured? Why are these pending tests not being run? | 13/01/2023 MF requested more recent sample id's so this can be further investigated by Trevor and Siemens. 20/01 SID's still needed to investigate. 27/01/23 Conor to trial a scenario next Thursday with Theopheline and investigate if pending tests remain when a sample is disposed and feedback. 03/02/23 Outcome of trial to be seen next week 24/02/23 Appears that samples with pending tests are being routed to IOM therefore not being disposed of - needs to be monitored | Michael/Trevor/Conor
19 |  | QC Database Maintenance | Martina has used the QC Database Maintenance procedure but it does not work. The timeframe doesn’t work & all data is deleted. | 13/01/23 Database backup and archieve working on Atellica. Desinstaion for achieved data required. Oliver suggested NAS drive and this is on the network. Investigate how to map drive to Atellica, T assistance required.. Other option is to use ADM, Peter suggested both (belts and braces). 90 days on ADM, 6 months on Atellica. Note, PHI code required from Cruinn/Siemens to access data. 20/01 NAS drvier installed & credentials sent to Trevor. In short term, data can ce copied onto USB & saved onto NAS drive. This needs to be configured to occur automatically in long term. 27/01/23 Conor did back up during week and worked fine but automated method needs ot be implemented. Conor to follow up with Oliver 03/02/23 Atellica backup configured, Neill mentioned looking at an alternative location for storage of the patient database backup 10/02/23 Neill to configre rules to amend patient database storage location. 03/03/23 Neill discussed this with Siemens & it has been advised not to change the database storage location & stick wtih ADM. It is also advised to disable USB access. 10/03/23 Archive work to be completed & location t be defined by RHM 24/03/23 To be added to IT list | Neill/Trevor/Oliver
20 | 2023-01-13 | Atellica Needle issue | Neill observed some needles being needle slightly bent approx. 2mm. | Neill following up with Siemens as seen on other sites also. Vac error also seen and maybe related. 20/01 Neill advised that this is a local issue & will be monitored. 27/01/23 Neill and Declan working on this. | Neill
21 | 2023-01-13 | Atellica Router upgrade | Neill to schedule router firmware upgrade on both Atellica lines. | 20/01 To be scheduled 27/01/23 Short line completed, long line to be scheduled. | Neill
22 | 2023-01-13 | Training | Atellica Maintenance Training | Olive trained Sorcha on Atellica Maintenance on line 1. Follow up training to be given to Emer next week. Off site training in Siemens is an option if required in future. 27/01/23 Olive off so to reschedule this training. | Olive
23 | 2023-01-13 | Open Channel | Olive sent open channel procedure sent to Martina. | 20/01 Edits to be made on procedure 27/01/23 SOP completed. Closed | Olive/Martina
24 | 2023-01-13 | UPS x3 | PM service 3 phase ups and Atellica ups x2 | Service being scheduled with Powervar. 20/01 Waiting on dates from Powervar 27/01/23 Scheduled for Feb 03/03/23 To be completed in next 2 weeks 24/03/23 PM's completed 31/03/23 UPS upgrade required for line 2 from 8KVa t 10 KVa. Ampage on current circuit breakwe needs to be increase to prevent trips. Cabelling also needs to be checked - all of these need to be completed for the IM system installation. - Fran to chase Shane Mulligan | Michael/Fran
25 | 2023-01-13 | Air compressor | PM service for Air Compressor | Service being scheduled with Atlas Copco 20/01 Waiting on dates from Atlas 27/01/23 Waiting on dates from Atlas | Michael
26 | 2023-01-13 | DI Water | MF to follow up with millipore re. second remote connection. | 20/01 Millipore to confirm 2nd connection; Martina noted an alarm that didn’t alarm outside 27/01/23 Paddy Dowling has scheduled to do this work. Currently two pump alarms showing, MF and Paul to follow up. Update from Ross Whitney Millipore, pumps ordered and will be replaced during pm which is being scheduled. 2nd remote connection will also be configured. 03/02/23 Graham from Millipore to carry out PM's next week & connect 2nd system to MiliQ Connect. 10/02/23 Graham carried out PM but not connection to MiliQ Connect. Pumps & membranes to be replaced. Should have been communicated to Graham before visit. MF 03/03/23 PM completed & 2nd system connected to MuliQ Connect but outstanding items to be completed - 2nd pump, flush RO membrane; May be a requirement to add pump to the mains - Milipore to discuss with faciliaties next time on site. 10/03/23 Paddy was on site this week - all outstanding work has been completed. Discussed pump requirements with Shane from maintenance. Agreed to install small pump & assess if this resolves the issue 24/03/23 Fran to follow up with Shane on pump installation as this is required for the additonal IM unit 31/03/23 Fran to follow up with Shane on pump installation | Michael/Fran
27 | 2023-01-13 | Decapper/ CM #2 | Intermittent decapper head down failures. | Samples causing error all coming from CM 2, MF adjusted z axis DW2 sensor. Issue now fixed. Closed | Michael
28 | 2023-01-13 | Desealer/Sealer #1 | Intermitteent desealing issue. | MF realigned desealer and service. Some foil still not being desealed correctly, all coming from sealer 1. MF to working on sealer 1 next week. 20/01 MF serviced sealer 1, PM completed on sealer 2. To be monitored. Martina noted issue w/ desealing sample post work 27/01/23 No complaints this week, Paul to monitor during off peak hours. | Michael
29 | End of 2022 | DMS | DMS clock out by 37 mins | 13/01/2023 Siemens x2 dialled in remotely and syrnchronised. Track and analysers had to be paused during this time. Synchronisation was successful. Closed | Michael/Siemens
30 | 2023-01-20 | Integrity Errors | Advised that there have been a number of sample integrity errors noted on the long line, patricularly on EQA material & paediatric samples. | Engineer to investigate 27/01/23 Still seeing these error, Declan following up with Siemens 03/02/23 Investigation ongoing 10/02/23 Declan advised that some work has been completed (air pump replaced, alignments etc) & the error rate will be monitored 03/03/23 Graham would like to understand the integrity error rate. We then need to investigate if the majority of errors are on front loaded samples such as paeds or EQA material or on patient samples as there is a different process for both. May be related to pumps & pressure adjustments? 10/03/23 Additional investigations to be completed. Ticket to be raised 24/03/23 Raised with Siemens Global - need to understand if it is possible to capture the # of integrity errors - add to IT list 31/03/23 Advised RHM that samples with integrity errors need to be visually inspected to determine issue. Global need to advise if it is possible to pull data on integrity errors to trend? | Neill/Declan
31 | 2023-01-20 | WASP Inventory Manager | To be implemented | 27/01/23 In progress, Sorcha and Trevor working on this. 10/02/23 Trevor has been working with Sorcha on this & progress has been made. Sorcha to train staff before go-live. Need to carry our stock check before go live & also manage 3rd party products. | Trevor/Sorcha
32 | 2023-01-27 | Atellica CH Field Safety Notice | Reagent Carryover from LDL Cholesterol Direct (DLDL), Total Protein II (TP), Triglycerides (Trig), and Triglycerides_2 (Trig_2) Impacting Magnesium (Mg) Results | 27/01/23 Graham Lee notified. Separting assays on multiple analysers and/ or batch testing causing difficulties for Lab. This issue is resolved by the addition of a Reagent Probe Cleaner 2 (RPC2) wash to the assays listed in Table 1 and will be implemented in Atellica Solution Software (SW) v1.27 which will be released soon. MF to find out when is v1.27 being released and can you go from 1.25 straight to 1.27 bypassing v1.26. Paul mentioned 3rd party reagents not affected, can these be used. Enda to follow up. 03/02/23 Assays being turned on & off at the moment - not ideal 10/02/23 Olive to investigate if setting up Trig-2 as open channel is a good option until SW update is avaiable. EC/PC to provide immediate update when possible on whether an interim patch is available or whether full SW upgrade to 1.27 is required to resolve the issue. 24/02/23 Olive working on setting up Mg as open channel in Temple St. Will do the same for Mullingar once successful 03/03/23 Open channel configuration & validation of magnesium completed by Olive - RHM happy with data. IT to be configured by Trevor 10/03/23 Issue with plasma samples running on long line. Will run if front loaded so likely to be configuration within ADM & DMS - Conor & Declan to check settings. | Michael/Enda/Emma/Peter
33 | 2023-02-24 | No result' tests | Regularly there are samples have 'no result' against certain tests & there is no way to determine the cause, such as integrity error, sampling issue etc. This is a software issue that needs to be addressed | Ticket to be raised with Siemens 03/03/23 Neill/Trevor to investigate - may improve with magnesium solution in place & lipids running ok 31/03/23 Ticket raised regarding TSH issue - samples not redirected to alternative analyser when other was in troubleshooting mode. This was tested this morning & sample was successfully rerouted to available system. SID's to be traced to determine root cause. | Olive/Enda/Neill/Michael/Trevor
34 | 2023-02-24 | Poor quality parts | RMH have concerns over the poor lamp quality, pump failures etc - it appears that many of the parts are of poor quality | Ticket to be raised with Siemens | Michael/Neill
35 | 2023-02-24 | Hilda/Justin Visit | No follow up from Hilda/Justin's visit in Sept - disappointed with the lack of follow up | Follow up call to be scheduled 03/03/23 Meeting to be scheduled in person with Siemens representatives. | Peter
36 | 2023-03-31 | GP samples after 5pm | Offline spinning of GP samples | Is it possible to confirgure a rule in ADM that any GP samples loaded after 5pm are spun only & not analysed? Alternative, can a location be defined on the IOM to load such samples? This will ensure there is no | Michael/Neill/Trevor
37 | 2023-03-31 | Atellica Tablets | Atellica Tablets not in use | Need to be added to the IT list | Peter

## Completed actions
Item # | Date Raised | Action Item | Description of issue | Comments/updates | Responsible | Status
3 | 2022-12-16 | Lithium Assay | 10 day TAT from Eurofins is not acceptable. Validation of Advia method is still ongoing. Cruinn will press Siemens for further information on the root cause of the issue. | All data generated has been sent to Paul/Martina for review. 13/01/23 Graham signed off on this and going live on Monday. 20/02 New test name (Xlith) to be added to SMS/Aptio server - MF to complete today. Closed | Emma/Olive/Martina
4 | 2022-12-16 | RSM Rack issue | Racks not being recognised & racks getting stuck – Emmett is on site 16/12 to troubleshoot this issue. | Monitor last week, no errors. Issue closed. Closed | Emmet/Michael
|  |  | Dilution should be completed if the dilution request is on the instrumet. On board dilution is sufficient for up to 15 photometric tests - unlikely to have >15 on UCFP samples. To be checked on CRP samples | Dilutions may need to be configured on Atellica rather than ADM to ensure on board dilution is used rather than needing sample to be retrieved from RSM. 13/01/23 Reference email from NOL sent Tue 10/01/2023 14:22 for details. Closed | Trevor/Olive
20 | 2023-01-13 | Atellica Needle issue | Neill observed some needles being needle slightly bent approx. 2mm. | Neill following up with Siemens as seen on other sites also. Vac error also seen and maybe related. 20/01 Neill advised that this is a local issue & will be monitored. 27/01/23 Neill and Declan working on this. | Neill
21 | 2023-01-13 | Atellica Router upgrade | Neill to schedule router firmware upgrade on both Atellica lines. | 20/01 To be scheduled 27/01/23 Short line completed, long line to be scheduled. | Neill
23 | 2023-01-13 | Open Channel | Olive sent open channel procedure sent to Martina. | 20/01 Edits to be made on procedure 27/01/23 SOP completed. Closed | Olive/Martina
27 | 2023-01-13 | Decapper/ CM #2 | Intermittent decapper head down failures. | Samples causing error all coming from CM 2, MF adjusted z axis DW2 sensor. Issue now fixed. Closed | Michael
28 | 2023-01-13 | Desealer/Sealer #1 | Intermitteent desealing issue. | MF realigned desealer and service. Some foil still not being desealed correctly, all coming from sealer 1. MF to working on sealer 1 next week. 20/01 MF serviced sealer 1, PM completed on sealer 2. To be monitored. Martina noted issue w/ desealing sample post work 27/01/23 No complaints this week, Paul to monitor during off peak hours. | Michael
29 | End of 2022 | DMS | DMS clock out by 37 mins | 13/01/2023 Siemens x2 dialled in remotely and syrnchronised. Track and analysers had to be paused during this time. Synchronisation was successful. Closed | Michael/Siemens
5 | 2022-12-16 | Dilution issue | CRP & UCFP automated dilutions are not working correctly. Discussion around whether these should be configured on Atellica Software rather than ADM to make appropriate use of the on-board dilution ring. Need to understand if this is an issue for all dilutions or specific to these 2 assays. | Work completed on diluton hardware on instruments - likely to have resolved issue - needs to be monitored 13/01/23 Neill worked on this during the week and its working as expected now. 20/01 May need to configure for Trig dilutions. HCG initial dilution causing delayfor EPU samples. - need to invesitgate if front loading EPU samples may resolve this 27/01/23 Neill working on analyser so unavailable, to get update next week 03/02/23 Paul to confirm if CRP & UCFP dilutions are working correctly. | Neill
1 | 2022-12-16 | LDL Performance issue | Concerns expressed by a number of GP’s that LDL results > Cholesterol. Because of INAB findings, LDL is not currently being reported & there is a calculation in place as an alternative. Graham has expressed that he feels clinically exposed because of this issue & requires a resolution as soon as possible. Currently reviewing correlation data between old & new method. From provisional data, it looks like the new method may not resolve the bias completely. Still approximately 5% bias but additional data needs to be generated to confirm. It would be helpful to review the validation data to understand how the Atellica LDL performance compared to the Beckman method. The updated Lipid assays (LDL, HDL & Trigs) all need to go live concurrently. This may involve configuring a rule on ADM where a calculated LDL may be required where the rules are violated. | Trevor & Conor to configure trigger for particular samples in ADM. Data for LDL assay, as well as other lipid assays, has been sent to Paul & Martina. Graham to review. Plan to resume running 'old' LDL method with rule in place until a decision is made based on the data. Looks like new method is approx 5-6% different from old method but to be confirmed. 13/01/23 Graham reviewing data, waiting sign off, possibly next week. Trevor and Conor have calculation working on ADM side. 20/01 Rule now in place for fasting & random lipids. If rule is breached (>15%), samples will be sent to the Mater for cholesterol testing. 27/01/23 Conor seeing 1 or 2 per day. Rule is working. Continue to monitor for the moment before closing out. 03/02/23 Plan to move to new assays next week. ADM rules to be moved to new method - TREVOR 10/02/23 Some results are being held in ADM & some in the LIS which is causing a delay in Clinicans seeing results. Trevor to investigate with Martina | Olive/Enda/Graham/Paul/Martina/Trevor
22 | 2023-01-13 | Training | Atellica Maintenance Training | Olive trained Sorcha on Atellica Maintenance on line 1. Follow up training to be given to Emer next week. Off site training in Siemens is an option if required in future. 27/01/23 Olive off so to reschedule this training. | Olive
6 | 2022-12-16 | EQA Testing | Need to configure the system to run EQA samples on a particular CH930 system on the long line. Neill initially hoped this could be configured through ADM but not successful. Is it possible to set up a panel for a particular instrument – Olive/Enda to investigate. | Olive spoke with UK FAS & this has been configured on ADM in Scottish site - Olive to follow up & get specific configuration details. 13/01/23 Neiling continue working on this, setup configuration and needs to test it. Instruction sent to Conor on how to run this. 20/01 - Configuration seems to be working well except when a reagent is missing from the particular system. Staff will need to know what assays are not available on board. 27/01/23 Lab to define configuration and test. 03/02/23 Lab testing to be completed 10/02/23 Update needed from Conor 24/02/23 Working effectively except if a reagent is missing on board, samples will be routed to different analyser | Lab
18 |  | Sample disposal | It seems that samples that have scheduled/rerun tests are being disposed of from the RSM. Can a rule be configured? Why are these pending tests not being run? | 13/01/2023 MF requested more recent sample id's so this can be further investigated by Trevor and Siemens. 20/01 SID's still needed to investigate. 27/01/23 Conor to trial a scenario next Thursday with Theopheline and investigate if pending tests remain when a sample is disposed and feedback. 03/02/23 Outcome of trial to be seen next week 24/02/23 Appears that samples with pending tests are being routed to IOM therefore not being disposed of - needs to be monitored | Michael/Trevor/Conor
25 | 2023-01-13 | Air compressor | PM service for Air Compressor | Service being scheduled with Atlas Copco 20/01 Waiting on dates from Atlas 27/01/23 Waiting on dates from Atlas | Michael
2 | 2022-12-16 | Batch testing - samples missed | Conor noted that x2 batches of LDL samples (x55 samples total) were front loaded onto the Sample Handler and on both occasions, only 18 results were produced – to be investigated by Cruinn | Olive advised that she successfully ran batches of lipid samples across all 3 platforms - To be monitored & tested. 13/01/23 still being monitored, OMG and Conor to do this next week. 20/01 Continues to be monitored - Conor to run dummy batch of samples when possible. 27/01/23 Olive to do this when next on site. 10/02/23 Olive to test when next on site 24/03/23 Batch testing completed & Olive confirmed that all samples were tested. Olive to run through batch testing procedure with RHM staff. | Olive/Conor
35 | 2023-02-24 | Hilda/Justin Visit | No follow up from Hilda/Justin's visit in Sept - disappointed with the lack of follow up | Follow up call to be scheduled 03/03/23 Meeting to be scheduled in person with Siemens representatives. | Peter

## VMM Issues
Serial No. | Time | Date | Issue | Resolution- if known | What else was happening on the Atellica at the Time | Logged with Cruinn/Siemens?